## Test set
Metric | YOLO26s 640 | YOLO26s 1024 | YOLO11s 1024 | How it is measured
SET-UP
Scans | 141 | 141 | 141 | CT scans in this split
Slices | 31655 | 31655 | 31655 | Every slice of every scan
Nodules (any reader) | 397 | 397 | 397 | Nodules marked by at least 1 of 4 radiologists
Nodules (>=2 readers) | 290 | 290 | 290 | Nodules marked by at least 2 radiologists (main target)
Confidence threshold | 0.4967 | 0.5201 | 0.555 | Chosen on validation: highest threshold still finding 90% of >=2-reader nodules; the same value is used on test
PER NODULE (whole scan)
CPM, >=2-reader nodules | 0.76 | 0.769 | 0.773 | Main score (LUNA16-style): average share of nodules found at 0.125, 0.25, 0.5, 1, 2, 4, 8 false alarms per scan
CPM, all nodules | 0.669 | 0.681 | 0.678 | Same, counting every nodule
Nodules found, >=2 readers | 87.9% | 89.7% | 87.6% | Per-nodule sensitivity at the chosen threshold
Nodules found, all | 79.3% (315/397) | 82.4% (327/397) | 78.6% (312/397) | Per-nodule sensitivity at the chosen threshold
False alarms per scan | 1.929 | 1.908 | 1.773 | False-alarm candidates (boxes linked across slices) per scan, at the chosen threshold
Found <6 mm (>=2 readers) | 58/76 | 61/76 | 58/76 | By nodule size
Found 6-10 mm (>=2 readers) | 112/120 | 115/120 | 113/120 | By nodule size
Found >=10 mm (>=2 readers) | 85/94 | 84/94 | 83/94 | By nodule size
Nodule-free scans with no false alarm | 10/20 | 8/20 | 9/20 | Scan-level specificity
Found at 0.125 false alarms/scan (>=2 readers) | 44.1% | 46.2% | 50.0% | FROC curve point
Found at 0.25 false alarms/scan (>=2 readers) | 57.2% | 58.6% | 62.1% | FROC curve point
Found at 0.5 false alarms/scan (>=2 readers) | 70.7% | 71.0% | 70.3% | FROC curve point
Found at 1 false alarms/scan (>=2 readers) | 80.0% | 83.5% | 81.7% | FROC curve point
Found at 2 false alarms/scan (>=2 readers) | 88.6% | 89.7% | 87.9% | FROC curve point
Found at 4 false alarms/scan (>=2 readers) | 94.8% | 93.1% | 93.5% | FROC curve point
Found at 8 false alarms/scan (>=2 readers) | 96.5% | 96.2% | 95.5% | FROC curve point
PER SLICE
True positive slices | 1227 | 1226 | 1197 | Nodule slice with a correct box
False negative slices | 805 | 806 | 835 | Nodule slice with no correct box
False positive slices | 519 | 492 | 424 | Nodule-free slice with a box
True negative slices | 29104 | 29131 | 29199 | Nodule-free slice with no box
Slice sensitivity | 60.4% | 60.3% | 58.9% | TP / (TP + FN)
Slice specificity | 98.2% | 98.3% | 98.6% | TN / (TN + FP)
Slice accuracy | 95.8% | 95.9% | 96.0% | (TP + TN) / all slices; high because ~94% of slices have no nodule
Slice balanced accuracy | 79.3% | 79.3% | 78.7% | Average of sensitivity and specificity
Slice precision | 70.3% | 71.4% | 73.8% | TP / (TP + FP)
Slice F1 | 0.65 | 0.654 | 0.655 | Balance of precision and sensitivity

## Validation set
Metric | YOLO26s 640 | YOLO26s 1024 | YOLO11s 1024 | How it is measured
SET-UP
Scans | 140 | 140 | 140 | CT scans in this split
Slices | 32734 | 32734 | 32734 | Every slice of every scan
Nodules (any reader) | 378 | 378 | 378 | Nodules marked by at least 1 of 4 radiologists
Nodules (>=2 readers) | 290 | 290 | 290 | Nodules marked by at least 2 radiologists (main target)
Confidence threshold | 0.4967 | 0.5201 | 0.555 | Chosen on validation: highest threshold still finding 90% of >=2-reader nodules; the same value is used on test
PER NODULE (whole scan)
CPM, >=2-reader nodules | 0.801 | 0.793 | 0.815 | Main score (LUNA16-style): average share of nodules found at 0.125, 0.25, 0.5, 1, 2, 4, 8 false alarms per scan
CPM, all nodules | 0.716 | 0.708 | 0.724 | Same, counting every nodule
Nodules found, >=2 readers | 90.0% | 90.0% | 90.0% | Per-nodule sensitivity at the chosen threshold
Nodules found, all | 83.1% (314/378) | 82.8% (313/378) | 81.8% (309/378) | Per-nodule sensitivity at the chosen threshold
False alarms per scan | 2.086 | 2.071 | 1.7 | False-alarm candidates (boxes linked across slices) per scan, at the chosen threshold
Found <6 mm (>=2 readers) | 47/60 | 48/60 | 46/60 | By nodule size
Found 6-10 mm (>=2 readers) | 129/142 | 129/142 | 130/142 | By nodule size
Found >=10 mm (>=2 readers) | 85/88 | 84/88 | 85/88 | By nodule size
Nodule-free scans with no false alarm | 10/20 | 9/20 | 10/20 | Scan-level specificity
Found at 0.125 false alarms/scan (>=2 readers) | 55.2% | 52.4% | 57.9% | FROC curve point
Found at 0.25 false alarms/scan (>=2 readers) | 64.5% | 64.5% | 66.9% | FROC curve point
Found at 0.5 false alarms/scan (>=2 readers) | 76.9% | 73.8% | 77.2% | FROC curve point
Found at 1 false alarms/scan (>=2 readers) | 84.5% | 85.2% | 86.2% | FROC curve point
Found at 2 false alarms/scan (>=2 readers) | 89.7% | 89.7% | 90.3% | FROC curve point
Found at 4 false alarms/scan (>=2 readers) | 93.8% | 92.4% | 94.5% | FROC curve point
Found at 8 false alarms/scan (>=2 readers) | 96.2% | 96.9% | 97.2% | FROC curve point
PER SLICE
True positive slices | 1372 | 1371 | 1314 | Nodule slice with a correct box
False negative slices | 690 | 691 | 748 | Nodule slice with no correct box
False positive slices | 527 | 518 | 383 | Nodule-free slice with a box
True negative slices | 30145 | 30154 | 30289 | Nodule-free slice with no box
Slice sensitivity | 66.5% | 66.5% | 63.7% | TP / (TP + FN)
Slice specificity | 98.3% | 98.3% | 98.8% | TN / (TN + FP)
Slice accuracy | 96.3% | 96.3% | 96.5% | (TP + TN) / all slices; high because ~94% of slices have no nodule
Slice balanced accuracy | 82.4% | 82.4% | 81.2% | Average of sensitivity and specificity
Slice precision | 72.2% | 72.6% | 77.4% | TP / (TP + FP)
Slice F1 | 0.693 | 0.694 | 0.699 | Balance of precision and sensitivity

## Notes
Per-nodule and per-slice results from 7eval_yolo.py, on every slice of every scan.
Boxes on neighbouring slices are linked into one candidate; a nodule counts as found if a candidate overlaps it.
The threshold is chosen on the validation set and reused unchanged on the test set, so the test set is the fair comparison.
YOLO's own scores (precision, recall, mAP) are in metrics_comparison_all_models.xlsx.
Source files: <model> results/results per nodule slice/<test¦val>/summary.json